Measurement goal

Pivot is a device that aims to improve speech skills by challenging users with balance exercises. The user has to maintain stability while standing on the Pivot and stability is measured based on the angle of the platform. The user strives to keep the platform horizontal throughout the sessions and a stability score is evaluated through checking how much and for how long the platform deviates from equilibrium (horizontal state). Therefore the tilt of the platform needs to be accurately tracked at all times in order to calculate the stability score and to control the movement of the actuators in both passive and active modes. Different methods and techniques of tracking orientation are considered and evaluated based on how well they suit are requirements.

Accelerometer
Accelerometer can measure acceleration on all three axis. It is used to determine the
direction of gravity so that it can be used as a reference for the roll and pitch operations;
however, it cannot distinguish between gravity and acceleration.

Gyroscope
The Gyroscope measures the rate of rotation of the sensor around the three axes. It offers
fast and accurate data on the rotation along axes; however, it is plagued with problems
relating to cumulative error, which results in inaccuracy when calculating the angle.

Magnetometer
Magnetotometer determines the strength of the external magnetic field, thus giving absolute
bearing information especially the yaw bearing. Nevertheless, it is highly susceptible to
external interference from motors and cables of the drones themselves.

Camera-based Visual Servoing
The camera-based approach to estimating the orientation relies on visual information
extracted from images for keeping track of the location and orientation of the system
compared to its surrounding. It uses downward or upward-facing cameras for tracking the tilt of the system relative to a reference. External cameras need to be precisely mounted and calibrated which would cancel the portability of the device, furthermore using visual feedback introduces complexity, cost, and sensitivity to the environment. Additionally, realizing visual servoing in real time requires special hardware support such as GPUs. Realization of visual servoing usually takes about 50–200 milliseconds per control iteration, which may exceed a fast control loop period needed for fast actuation.

IMU
An IMU combines a gyroscope, accelerometer, and optionally a magnetometer into a single
integrated package. Rather than relying on any one sensor alone, the IMU fuses the outputs
of all three to create a sensor fusion and with this sensor fusion, typically implemented
with a Kalman, Madwick or Mahony filter, to overcome the individual limitations of each
sensor to produce an reliable, continuous output of orientation across all three axes.

The IMU is proposed to be the sensor providing data to track the orientation of the
patform. The IMU itself outputs raw sensor measurements that do not inherently show
orientation, however those measurements can be used to determine distinct states of
orientations that can be represented as one of two possible forms.
The first method to represent attitude is with Euler angles, which is an intuitive way to
express orientation. The rotation around the z-axis will be called yaw and denoted with ψ,
the rotation around the y-axis will be referred to as pitch and denoted θ and the rotation
around the x-axis will be called roll and denoted with ϕ. Essentially, in the body frame,
yaw and ˙ψ (rate of change of yaw) measure the spin, whereas pitch and roll measure the
swing magnitude. However, Euler angles suffer from a problem called gimbal lock, which
is a situation where two rotation axes align and rotations become indistinguishable.
The second is the quaternion form, which is a mathematically robust way to describe
orientation. It isn’t an intuitive representation, just by looking at it the viewer lacks
physical intuition of what the attitude is, but every quaternion reliably and uniquely
denotes the corresponding orientation state. A single quaternion consists of a single scalar
(real) component and a three-dimensional vector (imaginary) component defined by the
complex basis vectors i, j, and k. It is usually written as:
q = w + x · i + y · j + z · k
The current quaternion of the platform can be found with sensor fusion, feeding the raw
IMU measurements to a state estimator algorithm like a Madgwick filter or a Kalman
based filter, and the state estimator will output the current quaternion derived from the
sensor measurement. With some further calculations the quaternion form can be converted
to Euler angles form and vice versa, but for algorithms, computer calculations, and more
complex systems the quaternion form is preferred as it is mathematically robust and does
not suffer from the problem of gimbal lock.
In project’s specific application, gimbal lock would normally not be a concern as the max tilt is mechanically limited to a considerably smaller amount than 90 degrees.

The device as well as all control systems require an accurate feedback about the controlled
system and this is provided by sensors. The measurements of one single sensor can rarely
be accurate and precise enough for the application (consistent orientation tracking). This
statement is true for the majority of engineering applications, as a single sensor will suffer
from noise, it is prone to drift and by itself it can rarely provide sufficiently accurate data.
Sensor fusion is the process of combining different sensor measurements, using sensors
that can complement each other, correct each other and ultimately provide data that is
accurate, precise and consistent enough for engineering systems, especially safety critical
devices like ours. Accelerometers have high-frequency noise, and gyroscopes suffer from
low-frequency drift. A state estimator is essentially a sensor fusion algorithm and in
our case it is used to combine the raw data of the accelerometer, magnetometer and the
gyroscope in a way that the different sensors complement and correct each other to provide
as reliable orientation estimate as possible.

Gyroscope Bias Drift
A fundamental limitation of MEMS gyroscopes is bias drift: a small constant offset in
the angular rate measurement that accumulates into a growing orientation error when
integrated over time. Cechowicz [ 1] characterises this as one of the primary error sources in
MEMS inertial navigation and demonstrates that without active compensation, integration
error grows continuously during operation. The Madgwick filter incorporates a bias drift
compensation term directly into its gradient descent formulation through the parameter
ζ, which estimates and subtracts the gyroscope bias at each update step [ 2]. This means
that over the duration of a rescue operation, the filter actively corrects for drift rather
than allowing it to accumulate unchecked.

Vibration Robustness
Helicopter airframes transmit strong broadband mechanical vibration through the fuselage
and winch cable. Lee and Hong [3] conducted flight tests on a real helicopter instrumented
with MEMS IMUs and documented large high-frequency noise on both the accelerometer
and gyroscope outputs induced by helicopter vibration. This vibration corrupts the raw
sensor readings and must be attenuated by the fusion filter before it reaches the controller,
unfiltered vibration aliased into the orientation estimate would cause the controller to
fire thrusters unnecessarily, wasting actuator authority. The complementary nature of
Madgwick’s algorithm, which integrates gyroscope data at high rate and corrects it slowly
via the accelerometer and magnetometer gradient, acts as an implicit low-pass filter on
the vibration content, producing a smooth orientation estimate even under significant
broadband excitation.

Magnetic Disturbance
Magnetometers are used in all three candidate filters to provide an absolute heading
reference for yaw. However, the helicopter environment presents a severe challenge: large
ferromagnetic structures, electrical motors and high-current wiring all distort the local
magnetic field significantly. Fan et al. [4] characterise the effect of such disturbances and
show that the RMSE of the heading (yaw) error caused by magnetic interference can reach
up to 15.4◦ in disturbed environments, compared to negligible error in magnetically clean
conditions. This is consistent with the findings of Ludwig and Burnham [ 8], who attribute
the poor yaw RMSE of all three filters in their experiments to magnetic disturbances present
in the data collection room. Wondosen et al. [5] further demonstrate on UAV data that
magnetic disturbances near the vehicle structure are a primary source of heading estimation
failure, and propose that algorithms should include active detection and rejection logic for
anomalous magnetometer readings. For our prototype, the magnetometer is calibrated to
partially compensate for the static hard-iron and soft-iron effects caused by the drone’s own
motors and wiring as described by Silic and Mohseni [ 6]. Residual dynamic disturbances,
for instance when the stretcher passes close to the helicopter winch motor, remain a
risk to yaw estimation accuracy, and this is acknowledged as a limitation of the current
implementation. In a production system, adaptive magnetometer weighting or a dedicated
disturbance rejection scheme would be required.

Figure 1: Sensor Fusion Approaches To Orientation Estimation

Various types of state estimator algorithms exist, the main archetypes are complementary,
gradient descent and Kalman-based sensor fusions. The Kalman-based sensor fusion
variations are considered to be the industry benchmarks [8] in terms of reliability and
comparing the performances of sensor fusion algorithms is an active field of research.
Ludwig and Burnham [8] compared all three archetypes directly on real quadcopter flight
data, and their findings confirm that all three approaches, Gradient descent (Madgwick),
Complementary (Mahony) and Kalman-based (EKF), are capable of providing accurate
orientation estimates in aerial vehicle applications.

Accuracy
Ludwig and Burnham [ 8] report RMSE values for roll, pitch and yaw of 3.47◦, 3.35◦ and
12.08◦ for Madgwick, and 4.12◦, 4.70◦ and 9.07◦ for Mahony respectively on quadcopter
flight data with Vicon ground truth. The EKF achieved the worst overall RMSE norm of
13.29◦ despite having the lowest individual roll and pitch errors (1.41◦ and 1.21◦), because
its yaw estimate was the most sensitive to the magnetic disturbances present during data
collection. All three filters therefore perform comparably in roll and pitch, while yaw
accuracy is dominated by the quality of the magnetometer signal rather than the choice of
fusion algorithm. An earlier comparison by Ludwig et al. [ 9] using foot-mounted MIMU
data with ground truth found Madgwick achieving RMSE values of only 0.59◦, 0.31◦ and
0.58◦ for roll, pitch and yaw when sensor noise is present, again outperforming the Mahony
filter across all three axes. Both studies confirm that Madgwick provides the best overall
orientation accuracy of the three candidate algorithms.

Execution Time and Embedded System Constraints
For our application the Madgwick filter was chosen, as it is a less complex algorithm
than the Kalman-based filters, while having approximately the same performance as EKF,
and slightly outperforming the Mahony filter. Execution time is a critical constraint in
our system: the control loop runs at 100 Hz, which leaves only 10 ms per cycle for all
processing including IMU read, sensor fusion, control computation and motor output.
On resource-constrained microcontrollers this budget is tight, and a computationally
heavy state estimator reduces the time available for the controller. Madgwick’s original
technical report [ 2] quantifies the computational cost of the filter at only 109 scalar
arithmetic operations per update for the IMU-only mode, and 277 operations when the
magnetometer is included. This very low operation count is a direct result of the gradient
descent formulation, which avoids the matrix inversions and iterative covariance updates
required by Kalman-based approaches. In measured runtime comparisons on the same
dataset of 6,401 samples, Ludwig and Burnham [ 8] report Madgwick completing in 0.208
s versus 0.290 s for the EKF, a 28% reduction in execution time. Mahony is slightly
faster still at 0.178 s, but as shown above, it yields lower accuracy. The significantly lower
mathematical complexity of Madgwick, compared to EKF, means faster execution time on
the microcontroller, preserving critical processing bandwidth for the controller and mixer
stages